## Slide 1
Developer: Pournaras Efthimios P., Student of Nuclear Physics - National & Kapodistrian University of Athens
Supervisor: Stiliaris Stathis, Assistant Professor of Physics, National & Kapodistrian University of Athens
IASA~magnetLAB group
Web page: http://www.iasa.gr/magnetLAB/magnetLAB.html
For info e-mail to: makis@iasa.gr
Introduction: Magnetic fields play a very important role in the science of physics (ex. particle accelerators), medicine (ex. NMR) and generally in every sector where modern technology has invaded. In many cases, the accurate knowledge of the magnetic field distribution in space is needed. The magnetLAB group has developed an automatic magnetic field mapping device, which can 3D-map the non-time-changing magnetic field of a selected area, by dividing it into several 2D planes. The output data of this device are used to design and to test special magnetic systems such as bending magnets for the beam transport of the 10 MeV CW-LINAC, which is also being developed at IASA.
The magnetic field mapping device:
The main aim of the magnetLAB group was to create a device which could scan a 2D accurately defined area. The whole process should be totally automated, so that after pre-scanning software adjustments no further external interferences should be needed. The reason for this is that accurate field measurements demand long time periods, which make it a boring and tiring process for the scientist who studies a magnetic system. The main parts of the mapping device are shown in figure 1. The device has the ability to measure fields in the areas of 0.3, 0.6, 1.2 and 3 Tesla. The error in measuring is 0.000005 Tesla and the field sensor (Hall Probe) takes 10 measurements per second. The relevant error for fields suitable for 5 MeV electron beams is 0.67/10000 and the half for the 10 MeV beams which will be developed in the future. The resolution in both axis is 1 twip which equals to 0.125 mm.
Figure 1:
"Diagram of the IASA - Magnetic Field Mapper"
Figure 2:
Automatic Magnetic Field Mapping Device’s Block Diagram
The twip unit has been introduced in order to refer to one full step of the stepping motors that are used to move the device's field sensor in a specific point of the 2D area that is mapped. As already said the device is fully controlled by a computer and when the mapping process starts no external interference is needed or aloud. The scientist can stop or pause the process. The software which has been specifically developed for the mapper has been sliced into two parts:
The data acquisition software (magnet1.exe) 2. The magnetic field reconstruction software (tran1.exe) The 1st part has been written for DOS in order to make use of the real time characteristic of this operating system. Although, it can run under Windows, too. It also makes use of the GP-IB scientific instruments communication technology. The 2nd part is written for Windows, because this operating system has a plethora of visualization tools. The magnetic field's data are stored in a ASCII format following a specific storing protocol, which has the recognizable extension MFM. The reconstruction software transforms the *.MFM files into 2D matrixes that can be visualized by Excel, MathCAD e.t.c. The MFM files contain data for magnetic field and temperature, that are extracted by Tran1.exe into *.MFF and *.MFT output ASCII data files.
Figure 3: "The scanning process"
In order to visualize the magnetic fields, which are being produced by the automated mapping device the group has developed RaceTRACK Version 1.0.0. An intergraded Windows® application, which can reconstruct and visualize the *.MFM, *.MFF magnetic field files. It can also visualize artificial fields, just for research. It also produces x-B(x) and y-B(y) section graphs, 2D Contour Plots and 3D Plots with full control over them. Furthermore it has the ability to emulate the orbit of a charged particle in the magnetic field distribution you currently process, using Lagrangian mechanics. Finally, it can calculate electron rigidities according to the inserted Kinetic Energy and find the required B field value for a given bending radius (for bending magnetic systems). The application is able to do several corrections to faulty measurements and process the field according to taste.
Figure 4:
““3D Plot of two natural linear magnets”
Figure 5: “The orbit of an electron in the B field of an accelerator’s bending magnet”
Figure 6: “ x - B(x) graph of the two natural linear magnets”
Magnetic Field Samples:
Figure 7:
“3D Plot of an accelerator’s bending electromagnet”
Figure 8:
“2D Contour Plot of 2 natural linear magnets side by side”
Figure 9:
“3D Plot of 2 natural ring shaped speaker magnets”
Figure 10:
“The mapper while mapping the bending magnet”